## Content
File content
Field | Description
Date | ! UTC date (for SMAPEx-1), Eastern Standard Time (for SMAPEx-2)
Time | ! UTC time (for SMAPEx-1), Eastern Standard Time (for SMAPEx-2)
Ave Temp 2.5cm (degC) | ! Unidata temperature sensors (ordered by increasing depths)
Ave Temp 5cm (degC)
Ave Temp 15cm (degC)
Ave Temp 40cm (degC)
Ave Rad (mA) | ! Thermalert sensor Raw output (only Thermalert sensors)
Thermalert T (deg C) | ! Calibrated TIR target temperature (only Thermalert sensors)
Ave Batt (volts) | ! logger battery voltage
Ave Voltage1 (mV) | ! 0-5cm Thetaprobe raw output
Soil moisture 1 (v/v) | ! 0-5cm calibrated Thetaprobe soil moisture
Ave Voltage2 (mV) | ! 20-25cm Thetaprobe raw output
Soil moisture 2 (v/v) | ! 20-25cm calibrated Thetaprobe soil moisture
Total Rain (mm) | ! Rainfall
Ave Voltage (mV) | ! Leaf wetness sensor Raw output
Ave LeafWet (% Wet) | ! Leaf wetness index (0/100)
Time stamp Apogee | ! Apogee sensors timestamp (only Apogee)
Apogee Tbody (deg C) | ! Calibrated Apogee sensor temperature (only Apogee)
Apogee Ttarget (deg C) | ! Calibrated target temperature (only Apogee)

## TIR calibration & issues
3 different type of sensors were used during SMAPEx-1 and 2, Apogee SI-121 and Thermalert TX (new and old version)
SMAPEx-1
station | TIR type | Comments
YS1 | Apogee | No data logged due to logging problem
YS2 | Thermaler TX new | The new version of the Thermalert TX was discovered to have a much larger operating range (up to 2000C) than the old version. The sensitivity (and hence precision) of the new instrument is therefore much lower than the old instrument.
YS3 | Apogee | No data logged due to logging problem
YS4 | Thermaler TX old | wrong lens transmissivity setting (set to 1, real value 0.65) might affect the calibrated data (TIR temperatures seem too low). The data are provided with the calibration equation applied: Ttarget [C] = (Rad (mA) - 8.5481)/0.146 and scaling to correct for emissivity based on Stephen-Botzmann: Ttarget [K] = {(1/emissivity)*(Ttarget[C] +273.15)^4}^0.25 where emissivity was set = 0.95 (literature value for grass) NO CORRECTION FOR LENS TRANSMISSIVITY HAS CURRENTLY BEN APPLIED TO THE DATA. NO FORMULA CURRENTLY EXIST FOR CORRECTION OF LENS TRANSMISSIVITY
SMAPEx-2
station | TIR type | Comments
YS1 | Thermaler TX old | ok
YS2 | Apogee | ok
YS3 | Apogee | ok
YS4 | Apogee | ok
Thermalert TX Calibration |  |  | Apogee SI-121 Calibration
The sensor was calibrated in laboratory using water as target. |  |  | The factory calibration coefficients were applied (http://www.apogeeinstruments.com/infraredradiometer/)
Calibration was performed on 11-01-2010 and 17-08-2010
For SMAPEx-1, the 11-01 regression was used, for SMAPEx-2 the 17-08 regression was used

## Thetaprobe calibration
Thetaprobe Calibration
the thetaprobe row output (mV) was converted to a soil moisture value
using factory calibration parameters for organic soils:
√ε=4.70V3-6.40V2+6.40V+1.07
SM=(√ε-a0)/a1
where
ε | soil dielectric constant
V | Thetaprobe output in Volt
SM | Volumetric soil moisture (m3/m3)
a0,a1 | factory calibration coefficients
1.6, 8.4 for mineral soils
1.3,7.7 for organic soils
The use of such parameters leads to errors in SM of the order
of ±0.05 (m3/m3)